**FILT’RAY 2G 2-month in-line filter**

For rinsing medical devices

Reference: **AT030734**

Specification description

AQUATOOLS anti-Legionella and anti-germ in-line filter.

Suitable for filtering water to supply equipment used in the treatment of medical devices (e.g. endoscopes).

Device: MBS.

Membrane type: tubular microfiltration.

Membrane: polyethylene.

Molecular weight cut-off: 0.1 μm nominal and 0.2 μm absolute sterilizing grade.

Qualified water volume per duration of use: 1,320 gallons (6,000 liters).

Filtration flow rate: 1.7 gal/min at 43.5 psi and 06. gal/min at 14.5 psi

(filter only, excludes any flow rate restriction in the faucet).

Maximum upstream pressure at the point-of-use: 72.5 psi.

Maximum lifespan: 2 months after installation.

Thermal shock resistance: 158°F for 30 minutes.

Chemical shock resistance: active chlorine levels of up to 100 ppm for 2 hours (cumulative total).

Sterile filter.

FILT’RAY 2G sterile disposable water filters are:

CE marked Class I Medical Devices in Europe that comply with directive 93/42/CEE;

Class II Medical Devices in the U.S.A. with FDA-510(K) clearance;

Class I Medical Devices in Canada that comply with SOR/98-282.

Box of 10. Each filter is individually wrapped and sterile.